## Data
IntegraGen SA | [10 empty] | Document Code: |  | FAB-ENR-035
Quality Management System | [10 empty] | Revision: |  | 0
miRpredX 31-3p results calculation spreadsheet
Changes tracking (History of changes)
Revision | Description |  |  | Date |  | Written by |  | Released by
0 | Initial version |  |  | 2017-09-05 |  | C. David |  | L. Ramon S. Roy
1 | Minor adjustments ; add of the mention "Use with Microsoft Excel" |  |  | 2017-10-02 |  | L. Ramon |  | C. David
Warning: To be used only for primary analysis with Microsoft Excel. The use of other software does not guarantee the functionality and safety of this spreadsheet. Refer to the Instructions for Use for a complete analysis of the results.
Information for Kit Utilized for Testing | [5 empty] | Insert data in green boxes only
Lot number
Expiration date
Number of kit use
Operator
Operator name
Thermocycler brand and reference
qPCR equipment brand and reference
Date of analysis
Step 1: Reverse transcription
Samples
Position
n° | Line | Column | Loc. | ID# code
S1
S2
S3
S4
1 | [6 empty] | Legend: |  | Presence of duplicates, please assign a unique position to each patient.
2
3
4
5
6
7
8
9
10
11
12
NTC-RT
NTC-qPCR
Plate layout
|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
| A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| B | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| C | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| D | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| E | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| F | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| G | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| H | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Checked by |  | 1: Date; 2: Signature
Step 2: Real Time PCR
Samples
miR-31-3p |  | 7: miR-calibrator
|  | Position | [7 empty] | Position
n° | Line | Column | Loc. | ID# code |  |  |  | n° | Line | Column | Loc. | ID# code
S1 |  | 3: -miR-31-3p; 7: S1; 11: -miR-calibrator
|  |  |  | -miR-31-3p | [7 empty] | -miR-calibrator
S2 |  | 3: -miR-31-3p; 7: S2; 11: -miR-calibrator
|  |  |  | -miR-31-3p | [7 empty] | -miR-calibrator
S3 |  | 3: -miR-31-3p; 7: S3; 11: -miR-calibrator
|  |  |  | -miR-31-3p | [7 empty] | -miR-calibrator
S4 |  | 3: -miR-31-3p; 7: S4; 11: -miR-calibrator
|  |  |  | -miR-31-3p | [7 empty] | -miR-calibrator
1 |  | 3: -miR-31-3p; 7: 1; 11: -miR-calibrator
|  |  |  | -miR-31-3p | [7 empty] | -miR-calibrator
2 |  | 3: -miR-31-3p; 7: 2; 11: -miR-calibrator
|  |  |  | -miR-31-3p | [7 empty] | -miR-calibrator
3 |  | 3: -miR-31-3p; 7: 3; 11: -miR-calibrator
|  |  |  | -miR-31-3p | [7 empty] | -miR-calibrator
4 |  | 3: -miR-31-3p; 7: 4; 11: -miR-calibrator
|  |  |  | -miR-31-3p | [7 empty] | -miR-calibrator
5 |  | 3: -miR-31-3p; 7: 5; 11: -miR-calibrator
|  |  |  | -miR-31-3p | [7 empty] | -miR-calibrator
6 |  | 3: -miR-31-3p; 7: 6; 11: -miR-calibrator
|  |  |  | -miR-31-3p | [7 empty] | -miR-calibrator
7 |  | 3: -miR-31-3p; 7: 7; 11: -miR-calibrator
|  |  |  | -miR-31-3p | [7 empty] | -miR-calibrator
8 |  | 3: -miR-31-3p; 7: 8; 11: -miR-calibrator
|  |  |  | -miR-31-3p | [7 empty] | -miR-calibrator
9 |  | 3: -miR-31-3p; 7: 9; 11: -miR-calibrator
|  |  |  | -miR-31-3p | [7 empty] | -miR-calibrator
10 |  | 3: -miR-31-3p; 7: 10; 11: -miR-calibrator
|  |  |  | -miR-31-3p | [7 empty] | -miR-calibrator
11 |  | 3: -miR-31-3p; 7: 11; 11: -miR-calibrator
|  |  |  | -miR-31-3p | [7 empty] | -miR-calibrator
12 |  | 3: -miR-31-3p; 7: 12; 11: -miR-calibrator
|  |  |  | -miR-31-3p | [7 empty] | -miR-calibrator
NTC-RT |  | 3: -miR-31-3p; 7: NTC-RT; 11: -miR-calibrator
|  |  |  | -miR-31-3p | [7 empty] | -miR-calibrator
NTC-qPCR |  | 3: -miR-31-3p; 7: NTC-qPCR; 11: -miR-calibrator
|  |  |  | -miR-31-3p | [7 empty] | -miR-calibrator
|  |  | Legend: |  | Presence of duplicates, please assign a unique position to each patient.
Plate Layout
|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
| A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| B | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| C | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| D | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| E | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| F | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| G | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| H | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Checked by |  | 1: Date; 2: Signature
The miRpredX 31-3p test kit has received IVD CE Mark approval and is commercially available in countries recognizing the CE Mark, or with applicable health authority registrations. The miRpredX 31-3p test kit is not available for sale in the United States.
IntegraGen SA
Génopôle Campus 1 – Bât. 8
5, Rue de Henri Desbruères
91000 Evry, FRANCE
For professional use only. |  | 9: support-miRpredX@integragen.com
© 2017 IntegraGen SA. All rights reserved. |  | 11: www.integragen.com

## Standard Results
IntegraGen SA | [9 empty] | Document Code: |  | FAB-ENR-035
Quality Management System | [9 empty] | Revision: |  | 0
miRpredX 31-3p results calculation spreadsheet
Instructions
Enter Ct of each standard, acceptance criteria are calculated automatically
CALIBRATION CURVE ANALYSIS: EVALUATION OF THE ACCEPTABILITY OF QUANTIFICATION
Insert data in green boxes only
| Standard | ng | Log (ng) | Ct miR-31-3p | Ct miR-31-3p Mean | SD | Ct miR-Cal | Ct miR-Cal Mean | SD | DCt | Std Err
| S1 | 100 | 2 |  | NA |  |  |  |  | NA
| S2 | 50 | 1.699 |  | NA |  |  |  |  | NA
| S3 | 5 | 0.69897 |  | NA |  |  |  |  | NA
| S4 | 0.5 | -0.30103 |  | NA |  |  |  |  | NA
NTC-RT
NTC-qPCR
[6 empty] | Acceptability Criteria |  |  | Obtained
|  | ∆Ct standard 2 |  |  |  | 2,8+/-0,5 |  |  | NA
|  | Slope miR-31-3p |  |  |  | [-2,92 ; -3,92] |  |  | #DIV/0!
|  | Slope miR-calibrator |  |  |  | [-2,92 ; -3,92] |  |  | #DIV/0!
|  | r2 miR-31-3p |  |  |  | > 0,95 |  |  | #DIV/0!
|  | r2 miR-calibrator |  |  |  | > 0,95 |  |  | #DIV/0!
|  | NTC-RT |  |  |  | Ct>39 or "Undetermined" |  |  | Hors Spécifications
|  | NTC-qPCR |  |  |  | Ct>39 or "Undetermined" |  |  | Hors Spécifications
| Checked by |  | Date | Signature
The miRpredX 31-3p test kit has received IVD CE Mark approval and is commercially available in countries recognizing the CE Mark, or with applicable health authority registrations. The miRpredX 31-3p test kit is not available for sale in the United States.
IntegraGen SA
Génopôle Campus 1 – Bât. 8
5, Rue de Henri Desbruères
91000 Evry, FRANCE
| For professional use only. | [7 empty] | support-miRpredX@integragen.com
| © 2017 IntegraGen SA. All rights reserved. | [9 empty] | www.integragen.com

## Samples Results
IntegraGen SA | [7 empty] | Document Code: |  | FAB-ENR-035
Quality Management System | [7 empty] | Revision: |  | 0
miRpredX 31-3p results calculation spreadsheet
Instructions
Enter the Ct of each sample, the table calculates the status of each sample.
Analysis factor related to the standard |  | 1.43
Insert data in green boxes only
RNA ID | Ct miR-31-3p | Ct miR-31-3p Mean | SD | Ct miR-Cal | Ct miR-Cal Mean | SD | DCt | DDCt | Fold | miR-31-3p status
WookieeStephaneRoyLucasRamon
FAIBLE
WookieeStephaneRoyLucasRamon
FAIBLE
WookieeStephaneRoyLucasRamon
FAIBLE
WookieeStephaneRoyLucasRamon
FAIBLE
WookieeStephaneRoyLucasRamon
FAIBLE
WookieeStephaneRoyLucasRamon
FAIBLE
WookieeStephaneRoyLucasRamon
FAIBLE
WookieeStephaneRoyLucasRamon
FAIBLE
WookieeStephaneRoyLucasRamon
FAIBLE
WookieeStephaneRoyLucasRamon
FAIBLE
WookieeStephaneRoyLucasRamon
FAIBLE
WookieeStephaneRoyLucasRamon
FAIBLE
Mean miR-cal legends column: |  |  | Calibrator above alert limit: Result is acceptable under conditions* |  |  |  | Calibrator above acceptance limit low RNA content of sample: check sample RNA concentration/quality*
* See Instruction For Use
Checked by | Date | Signature
The miRpredX 31-3p test kit has received IVD CE Mark approval and is commercially available in countries recognizing the CE Mark, or with applicable health authority registrations. The miRpredX 31-3p test kit is not available for sale in the United States.
IntegraGen SA
Génopôle Campus 1 – Bât. 8
5, Rue de Henri Desbruères
91000 Evry, FRANCE
For professional use only. | [7 empty] | support-miRpredX@integragen.com
© 2017 IntegraGen SA. All rights reserved. | [9 empty] | www.integragen.com

## References (hidden)
Line | Column
A | 1
B | 2
C | 3
D | 4
E | 5
F | 6
G | 7
H | 8
9
10
11
12
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
A | A1 | A2 | A3 | A4 | A5 | A6 | A7 | A8 | A9 | A10 | A11 | A12
B | B1 | B2 | B3 | B4 | B5 | B6 | B7 | B8 | B9 | B10 | B11 | B12
C | C1 | C2 | C3 | C4 | C5 | C6 | C7 | C8 | C9 | C10 | C11 | C12
D | D1 | D2 | D3 | D4 | D5 | D6 | D7 | D8 | D9 | D10 | D11 | D12
E | E1 | E2 | E3 | E4 | E5 | E6 | E7 | E8 | E9 | E10 | E11 | E12
F | F1 | F2 | F3 | F4 | F5 | F6 | F7 | F8 | F9 | F10 | F11 | F12
G | G1 | G2 | G3 | G4 | G5 | G6 | G7 | G8 | G9 | G10 | G11 | G12
H | H1 | H2 | H3 | H4 | H5 | H6 | H7 | H8 | H9 | H10 | H11 | H12